| General Information |  |  |  |  |
| --- | --- | --- | --- | --- |
| Device Name & Number | OS01-99: OncoSil ™ Calibration System |  | Facility / Hospital Name |  |
| Device Batch Number |  |  | Ion Chamber Identification (e.g.: Make, Model and Location) |  |
| Date of Receipt / Delivery (dd/mmm/yyyy) |  |  |  |  |
| Activity Quoted on the Calibration Certificate (XXX.X MBq at 12:00 CET/CEST). Attach a copy to this form |  |  | Date and Time of this Calibration (dd/mmm/yyyy, hh:mm – 24 hours) |  |
| Reference Date on the Certificate (dd/mmm/yyyy) |  |  | OncoSil™ Preparation Start and Finish Times (hh:mm- 24hrs) |  |
| Is this a recalibration (if yes write the current ion chamber factor here) | YES ☐ NO ☐ |  | OncoSil™ Preparation Total Time (Start - Finish) |  |

| Determination of the Calibration Dose Activity at the Time of Measurement |  |  |  |  |
| --- | --- | --- | --- | --- |
| - Time Zone for Measurement Location (e.g. EDT or EDST) |  |  | - Therefore Time Difference (Days/ Hours) between Local Reference Time and Local Time of Measurement |  |
| - Relative Time Difference (hours) Currently between the Local Time and CET (CEST) |  |  | - Decay Factor Obtained from Tables for the Time difference between Local Reference and Time of Measurement |  |
| - Local Time Equivalent of Reference Time (CET/CEST + Relative Time Difference) |  |  | - Corrected Calibration Certificate Activity that the ion chamber will be adjusted to Reached (Calibration certificate Activity x Decay Factor) | _______________MBq or _____________mCi (see factor tables for conversion) |

| Record of Measurements and Calculations for Initial Calibration Dose Dilution – Ionisation Chamber Calibration |  |  |  |  |
| --- | --- | --- | --- | --- |
| Background (If applicable) | Background Corrected Measurement 1 (4 significant figure MBq or mCi) | Background Corrected Measurement 2 (4 significant figure MBq or mCi) | Background Corrected Measurement 3 (4 significant figure MBq or mCi) | Average Background Corrected Measurement (4 significant figure MBq or mCi) |
| Calculate the Relative Difference between the Average Measured Value and the Decay Corrected Certificate Value using |  |  | Note: % Diff < ±5%: ion chamber adjustment is not required. % Diff > ±5%: ion chamber factor should be adjusted. |  |
| % Diff=Average measured-Corrected CertificateCorrected Certificate×100 |  |  |  |  |

| Ion Chamber Calibration Factor Record |  |  |  |  |
| --- | --- | --- | --- | --- |
| 32P Calibration Factor before Reading OncoSil™ Calibration Suspension |  |  | Adjusted 32P Calibration Factor after Reading OncoSil™ Calibration Suspension |  |

| Confirmation of Calibration Measure |  |  |
| --- | --- | --- |
| Confirmation Activity Measurement – Record the value of the post-adjustment measurement. |  | If the Calibration was unsuccessful, please contact OncoSil Medical |
| Is the measured value within ±5% of the expected value per certificate? | ☐ YES Proceed with next step ☐ NO Contact OncoSil Medical |  |

| Second Dilution - Confirmation of the Time of Measurement (NOTE: these steps are optional for a recalibration exercise |  |  |  |  |
| --- | --- | --- | --- | --- |
| Time (hh/mm) Diluted Dose Measured |  |  | Is the time of the Diluted Vial measurement withing 1 hour of the Calibration Vial measurement | Yes/No. If No get a New Factor (NF) - see note NF = |

Note: If the measurement of the diluted vial is within 1 hour of the calibration vial the same correction factor can be used.

If the time is extended past 1 hour a decay factor from Tables in page 3 & 4 will be required to correct the Confirmation Activity Measurement

| Record of Measurements and Calculations for Vial 2 – Ion Chamber Calibration |  |  |  |  |
| --- | --- | --- | --- | --- |
| Background (If applicable) | Background Corrected Measurement 1 (4 significant figure MBq or mCi) | Background Corrected Measurement 2 (4 significant figure MBq or mCi) | Background Corrected Measurement 3 (4 significant figure MBq or mCi) | Average Background corrected Measurement (4 significant figure MBq or mCi) |

| Confirmation of Linearity of Calibration |  |  |  |  |
| --- | --- | --- | --- | --- |
| Calculate the Expected diluted value ExpectedActivityMBq/mCi=Confirmation ActivityMeasurementMBq/mCi×1.77×NFif applicable |  |  | Is the Average Corrected Measurement within ±10% of the expected diluted value? | ☐ YES Calibration completed ☐ NO Contact OncoSil Medical |

| Conducted By: |  |  |  |
| --- | --- | --- | --- |
|  | Name (BLOCK LETTERS) | Signature | Date (DD MM YYYY) |
| Checked By: |  |  |  |
|  | Name (BLOCK LETTERS) | Signature | Date (DD MM YYYY) |

[Once complete, forward this form to OncoSil Medical customer service by email (orders@oncosil.com.au) or fax (+61 2 9252 3988)]

NOTE:- to convert the activity in MBq to mCi divide all MBq figures by 37. Therefore 250 MBq = 6.757 mCi.

NOTE:- to convert the activity in MBq to mCi divide all MBq figures by 37. Therefore 250 MBq = 6.757 mCi.